# EEG Quality Assurance Checklists

# Checklist: Data Collection QA

## Equipment Setup

- Are all electrodes and cables inspected for damage (fraying, bent pins, worn contacts)?

- Is the cap sized and positioned correctly against standard landmarks (e.g., 10-20 system)?

- Are reference and ground electrodes placed and secure?

- Have all electrode impedances been checked and brought within the amplifier's recommended range?

- Are EOG, ECG, or EMG reference channels included if needed for artifact correction later?

- Is the sampling rate and hardware filter setting appropriate and documented?

- Is cable routing secured to minimize movement and swing artifacts?

- Is the recording environment free of avoidable electrical interference (monitors, chargers, unshielded cords nearby)?

- Has a short test recording confirmed all channels are active and noise-free?

- Has the participant been briefed on minimizing movement, jaw tension, and unnecessary blinking?

## Live Data Collection

- Is the raw trace being visually monitored during recording, not just logged blindly?

- Are event markers/triggers confirmed to be logging correctly and time-locked?

- Is impedance spot-checked during longer sessions?

- Are notable events (coughing, movement, environmental disruption) being logged in real time?

- Is there a plan for addressing excessive artifacts mid-session (re-gel, reposition, pause)?

- Are breaks scheduled to reduce fatigue-driven artifacts (drowsiness, muscle tension)?

- Is the recording being saved continuously with a backup in place?